I&C Service Engineer, Ore Enrichment Plant Madneuli

Svintsov Stanislav

28 September 2025

TECHNICAL SPECIFICATION
For the procurement of an Electromagnetic Flowmeter

| Item № | Parameter Name | Value |  |
| --- | --- | --- | --- |
| 1 | Customer | I&C Service, Ore Enrichment Plant Madneuli |  |
| 2 | Full name, position, phone, e-mail of the responsible executor / Customer’s representative | Svintsov S.P.., Engineer of I&C Service |  |
|  |  | Engineer of I&C Service |  |
|  |  | +995591957961 |  |
|  |  | SSvintsov@richmetalsgroup.com |  |
| 3 | Basis for procurement | Notice of unsuitability for use |  |
| 4 | Equipment name | Electromagnetic Flowmeter |  |
| 5 | Quantity to be procured | 1 set |  |
| 6 | Technical characteristics, requirements | Measured medium characteristics: |  |
|  |  | Name | Slurry (Pulpa) |
|  |  | Temperature, Со | +30…+45 |
|  |  | Excess pressure, kPa | 100 to 101,3 kPa |
|  |  | Ambient characteristics: |  |
|  |  | Temperature, Со | +10…+35 |
|  |  | Humidity, % | 60…80 |
|  |  | Aggressiveness | yes |
|  |  | Dust presence | yes |
|  |  | Flowmeter characteristics: |  |
|  |  | Transducer type | Electromagnetic |
|  |  | Process connection: | PN16, St37-2/FE 410W B, flange EN1092-1 (DIN2501) |
|  |  | Materials | Lining: polyurethane, hard rubber, PFA, PTFE, natural rubber Electrodes: 1.4435/304L, Alloy C-22, tantalum, platinum, Duplex 1.4465, tungsten carbide, titanium |
|  |  | Measured parameters | Volumetric flow, conductivity, mass flow |
|  |  | Measuring range: | 0...9'600 m³/h |
|  |  | Product temperature range/Ambient temperature range | -40...+180°C –40...+60 °C (–40...+140 °F) |
|  |  | Protection rating | IP 67 (Nema 4x) |
|  |  | Display | yes |
|  |  | Output signals | 4...20 mA (active/passive) Pulse-frequency (active/passive) Two status signal outputs |
|  |  | Digital communication | HART PROFIBUS PA FOUNDATION Fieldbus |
|  |  | Power supply / Display | 85–260 V AC, SEA, 4-wire + touch control, SEA = languages JA+ID+EN |
| 7 | Documentation | - Availability of operation and maintenance manual (documentation language – Russian, English). - Availability of type approval certificate for measuring instruments and calibration procedure – Not required. - Availability of verification certificate (initially verified) – Not required. |  |
| 8 | Warranty period | Not less than 1 year |  |
| 9 | Additional conditions | None |  |
| 10 | Other requirements | None |  |

Hide details

Liner- natural rubber

Process Connection-PN16, St37-2/FE 410W B, flange EN1092-1 (DIN2501)

Electrodes - 1.4310/302, brush type, meas. Electrode only

Calibration - 0.5%

Additional Test, certificate - Mat. certificate (tube + flange)

Approval - Non-hazardous area

Housing - Wall Alu, IP67 NEMA4X, -40oC/oFSensor see specification

Cable, Remote Version - ..... m coil and signal cable

Cable Entry - Gland M20 (EEx d > thread M20)

Power Supply; Display - 85-260VAC, SEA, 4-line + Touch control SEA= language JA+ID+EN

Adjustment; Software Feature - Factory setup; basic version

Output, Input - 4-20mA HART + freq. + 2x relay

Marking - Tagging (TAG), see additional spec.