## Sheet1
დასახელება | რაოდენობა | მოთხოვნა კანონმდებლობით | სასურველი მახასიათებელი | სავარაუდო მოდელი | ზონა | ღირებულება | საგარანტიო ვადა | მოწოდების ვადა | გადახდის პირობა | შენიშვნა/კომენტარი
ბიოუსაფრთხოების კარადა | 1 | (პჯრ- ბოქსი ან ლამინარული კარადა),რომელსაც გააჩნია ულტრაიისფერი გამოსხივების წყარო და ჰაერის ნაკადის ცირკულირების დანადგარი | -for particles of 0.5µm and more ISO 5 -for particles of 5.0µm and more ISO M (20; ≥5µm); LSAPC Cabinet class according to ЕN 12469, NSF/ANSI 49 II Cabinet type according to NSF/ANSI 49 A2 Class of the installed НЕРА-filters according to EN 1822-1 H14 Average velocity of the inflow though the work opening,m/s 0,47±0,03 Average downflow velocity in the working chamber, m/s 0,35+0,01 Clean air inflow volume, m3/h 1008-1036 Air outflow volume, m3/h 426-484 Air recirculation rate in the cabinet, % ≈70 Illuminance level in the working zone, lux, not less than, lx 2000 Noise level at 1m distance from the cabinet 53* dB | Lamsystems SAVVY BMB-II-Laminar-S-1,5 SAVVY /Microbiological Safety Cabinet Class II Type A2 ან Microbiological safety cabinet BMB-II-“Laminar S.”-1,2 (221.120) NEOTERIC Class II, type A2 Article 1Е-В.001-12.0 | ექსტრაქცია
ულტრაცენტრიფუგა (მაცივრიანი) | 1 | (ბრუნვის სიჩქარე არანაკლებ 6000 ბრ/წუთში) | Speed range 15,000rpm (200-15,000rpm) in 100rpm increments. Rotor capacity1.5/2mL ×24 ¦ 0.2/0.5mL ×36 ¦ PCR-8 strips×4 ¦ 5mL ×12 ¦ 5mL×18, refrigerated centrifuge, Temperature range -20℃ to 40℃ ან Max. Speed 15,000rpm (200-15,000rpm) in 100rpm increments. Max. RCF 21,380×g, step: 10×g Capacity 36 0.2/0.5mL, 24×1.5/2mL, 4×0.2mL PCR 8-tube strip, 12×5mL (cone bottom), 18×5mL (round bottom) Temperature setting -20°C-40°C Run Time 30s-99 min; HOLD (continuous running) Acceleration and deceleration time 25s↑ 25s↓ Motor Type Brushless DC motor Safety Performance Dual door lock, over-speed detection, over-temperature detection with Internal diagnosis system. Power Requirements Single-phase, 100-120V/200-240V, 50Hz/60Hz.500W Noise Level <56dB Dimensions (WxDxH] 332 × 553 × 283mm Weight 30kg Other functions Rotation speed/RCF switchover,inching operation, running processindicator,sound alert.Adjustable acceleration/deceleration at 9 levels program storage capacity. | TGL16 ან D1524R | ექსტრაქცია
თერმოშეიკერი (თერმობლოკი) | 1 |  | Temperature setting range +4°C ... +100°C Temperature control range 15°C below ambient ... +100°C Temperature setting resolution 0.1°C Temperature stability ±0.1°C Temperature accuracy at +37°C ±0.5°C Average heating speed from +25°C to +100°C 5°C/min Average cooling speed from +100°C to +25°C 5°C/min Average cooling speed from +25°C to +4°C 1.8°C/min Temperature uniformity over the block at +4°C ±0.6°C Temperature uniformity over the block at +37°C ±0.1°C Temperature uniformity over the block at +100°C ±0.3°C Temperature calibration coefficient range 0.936...1.063 (± 0.063) Speed control range 250–1400 rpm Digital time setting 1 min–96 hrs / non–stop (increment 1 min) Timer sound signal + Orbit 2 mm Display LCD, 16 x 2 signs Microprocessor controlled temperature, mixing speed and operation time + Maximum continuous operation time 168 h Overall dimensions (W×D×H) 220x240x90 mm Weight 3.7 kg Input current/power consumption 12 V, 4.9 A / 60 W External power supply Input AC 100–240 V; 50/60 Hz; Output DC 12 V | TS-100C | ექსტრაქცია
ვორტექსი (სანჯღრეველა) | 2 |  | Variable speed of 0 to 3,000rpm for gentle shaking to rapid mixing of samples Touch or continuous operation mode Includes a standard top head for mixing a single test tube or other small vessel up to 30mm in diameter Electronic speed control maintains constant speed even during changes in load Steel base with specially designed vacuum suction feet provides stability and dampens vibration Robust aluminum-cast construction, includes rectangle PCR tuve adapter | MX-S ან Thermo Scientific 88882010 ან HumaTwist(110 V/60 Hz) | ექსტრაქცია + შერევა
მინი ცენტრიფუგა (სპინერი) | 1 |  | Easy to operate with dual start/stop function — Close lid for a spin speed up to 7000 rpm Smooth and quiet operation (noise level ≤45dB) Accelerates and brakes in seconds Safe, endurable and robust rotor with clamp locking design Compatible with 2mLx8 and PCR8x4 rotors Quick-change rotor without any tool | D1008 ან HuMax Micro(110 - 230 VAC EU plug) | ექსტრაქცია
პიპეტები | 10 | პიპეტები ან ვარიაბელური პიპეტი | პიპეტები 2-20 მიუმოლამდე 2 ცალი, პიპეტები 10-100 მიუმოლამდე 1 ცალი, პიპეტები 1-10 მიუმოლამდე 1 ცალი, პიპეტები 20-200 მიუმოლამდე2 ცალი, პიპეტები 100-1000 მიუმოლამდე 2 ცალი, პიპეტები 10-100 მიუმოლამდე 1 ცალი, პიპეტები 5-50 მიუმოლამდე 1 ცალი | Thermo fisher, pipet4upro | ექსტრაქცია + შერევა
პიპეტების სადგამი | 3 |  | მარტივად წმენდადი
ბიოუსაფრთხოების ბოქსი | 2 | დნმ/რნმ- ზე სამუშაოდ განკუთვნილი ბიოუსაფრთხოების ბოქსი (საკმარისია ე.წ. მკვდარი (უძრავი) სივრცის უზრუნველყოფა) | Installation work chamber air cleanliness class for suspended particle (aerosol) concentration as per ISO 14644-1 -for particles of 0.5µm and more ISO 5 -for particles of 5.0µm and more ISO M (20; ≥5µm); LSAPC Class of the installed in the cabinet HEPA-filters according to EN 1822-1 H14 Prefilter according to EN 779 G4 Average downflow velocity in the work chamber: - preset velocity, mps 0,4 - recommended velocity range for customized setting that guarantees uniformity (laminarity) of the airflow, mps 0,25-0,5 Illuminance level in the working zone, lux, not less than 1000 Noise level at 1m distance from the cabinet (at centerline of the work openingwith the front window sash at working position), dB, not more than 55** | Laminar flow cabinet clean bench BAVnp-01-Laminar-S-1.8 LORICA ან Vertical Laminar Flow Cabinet BAVnp-01-“Laminar-S.”-1,2 LORICA Article 1Е-D.001-12.0 | შერევა
მინი ცენტრიფუგა (სპინერი) ფლეიტებისთვის და სტრიპებისთვის ვორტექსის ფუქნციით | 1 |  | SMS-cycle regulation 1 - 10 cycles Speed regulation range 500–3500 rpm Max. RCF 610 x g Setting resolution 100 rpm Display LCD, 2 x 16 signs Centrifugation mode time range 0–10 min (increment 1 s; after 1 min - 1 min) Vortex mode time range 0–5 min (increment 1 s; after 1 min - 1 min) Overall dimensions (W×D×H) 165 × 220 × 230 mm Weight 2.7 kg Input current/power consumption 12 V, 2 A / 24 W External power supply Input AC 100–240 V; 50/60 Hz; Output DC 12 V PCR tube strips using optional SH-12x8 adapter | MSC-2P | შერევა
ოპტიკური ბლოკი - რეალურ-დროითი პჯრ- ისთვის | 1 |  | Streamlined Smaller foot print: 27 W x 50 D x 40 H cm, designed to maximize horizontal and vertical bench space Front-facing fully motorized blockInteractive Touch screen interface enables viewing of amplification plots Easy and intuitive interface Bright light LED median lifetime greater than 5 years Less frequent maintenance 10 GB on-board memory for storage of 2,000–5,000 runs HID-validated systeმ 0.2 mL standard 96-well format Multiplexing capability for up to six dyes, enabling future quantification technolog Precise temperature control with Peltier-controlled blocks (known as VeriFlex zones USB interface Real-time PCR is a flexible technique that can be used for multiple forensic applications: Forensic databas--Laboratories without direct amplification workflows Forensic casework --Confirmatory screening of male DNA --Sexual assault sample workflow --Quality and quantity sample assessmentღია სისტემა სხვადასხვა წარმოების რეაგენტებისა და ექსტრაქციის ნაკრებებისთვის | QuantStudio™ 5 Real-Time PCR System, 96-well, 0.2 mL | ამპლიფიკაცია